Phlebology, Proctology, Gynecology, ENT, Neurosurgery, Dermatology, Liposuction, Therapy

|  |  |
| --- | --- |
| Phlebology | |
| 1、EVLA Thigh | 1470nm, CW 8w 80J/cm |
| 2、EVLA calf | 1470nm, CW 6w 60J/cm |
| 3、EVLA Small Vessel | 1470nm, CW 3w 30J/cm |

|  |  |
| --- | --- |
| Proctology | |
| 1、Hemorrhoids Grade 2 | 1470nm, CW 6w, 250 Joules per hemorrhoid |
| 2、Hemorrhoids Grade 3 | 1470nm, CW 7w, 350 Joules per hemorrhoid |
| 3、Hemorrhoids Grade 4 | 1470nm, CW 8w, 350 Joules per hemorrhoid |
| 4、Fistula | 1470nm, CW 10w 100J/cm |
| 5、Pilonidal Sinus | 1470nm, CW 10w 100J/cm |
| 6、Fissure Ablation | 1470nm, CW 7w 100J/cm |

|  |  |
| --- | --- |
| Cosmo gynecology | |
| 1、Laser Vaginal Tightening 1st Session 600 Joules  （Laser Vaginal Tightening 1st 600J） | 1470nm, Pulse 6w 500ms Ton, 500ms Toff |
| 2、Laser Vaginal Tightening 2nd Session 600 Joules | 1470nm, Pulse 7w 500ms Ton, 500ms Toff |
| 3、Laser Vaginal Tightening 3rd Session 600 Joules | 1470nm, Pulse 8w 500ms Ton, 500ms Toff |
| 4、Laser Vaginal Tightening 4th Session 600 Joules | 1470nm, Pulse 8w 500ms Ton, 500ms Toff |
| 5、Laser Vaginal Rejuvenation, 4 passes, total 360 Joules（Laser Vaginal Tightening 360J） | 1470nm, CW 6w 15s Ton |
| 6、Laser Vaginal Rejuvenation Plus, 4 passes, 600Joules | 1470nm, CW 10w 15s Ton |
| 7、Stress Incontinence, 75 Joules | 1470nm, CW 5w 15s Ton |
| 8、Infection control, 50 Joules | 1470nm, CW 5w 10s Ton |
| 9、Treatment after Menopause, 4 passes, total 480 Joules | 1470nm, CW 8w 15s Ton |
| Gynecology Hysteroscopy | |
| 1、Myomas | 1470nm 12w CW |
| 2、Polyps | 1470nm 8w CW |
| 3、Septum | 1470nm 10w CW |
| 4、Adhesion | 1470nm 10w CW |
| 5、Dysmorphie Uterus | 1470nm 12w CW |
| 6、Isthmocele | 1470nm 12w CW |
| 7、T-Shape Uterus | 1470nm 12w CW |
| Gynecology Laparoscopy | |
| 1、Myomas | 1470nm 12w CW |
| 2、Ovarian Cysts | 1470nm 10w CW |
| 3、Adhesion | 1470nm 10w CW |
| 4、Salpingectomy | 1470nm 10w CW |
| 5、Cysts | 1470nm 10w CW |
| 6、Endometriosis | 1470nm 10w CW |

|  |  |
| --- | --- |
| ENT | |
| 1、Vocal Polyps | 980nm 13w CW |
| 2、Cordectomy | 980nm 12w CW |
| 3、Laryngeal Carcinoma | 980nm 14w CW |
| 4、Haematomas | 980nm 14w CW |
| 5、Fibromas | 980nm 14w CW |
| 6、Larynxpapillomatosis | 980nm 14w CW |
| 7、Paracentesis | 980nm 13w CW |
| 8、Acoustic Neuroma | 980nm 13w CW |
| 9、Cholesteatoma | 980nm 10w CW |
| 10、Tonsillectomy | 980nm 13w CW |
| 11、Tumour Vaporisation | 980nm 13w CW |
| 12、Uvulopalatoplasty (LAUP) | 980nm 10w CW |
| 13、Conchotomy | 980nm 13w CW |
| 14、Dacryocystorhinostomy (DCR) | 980nm 10w Pulse, 50ms Ton, 90ms Toff |
| 15、Turbinate Reduction | 980nm 10w CW |
| 16、Nasal Polypectomy | 980nm 12w CW |

|  |  |
| --- | --- |
| Neurosurgery | |
| 1、PLDD L2-L3, L3-L4, L5-S1 Total 800-1500J | 1470nm, Pulse 7w 1s Ton, 5s Toff |
| 2、PLDD L4-L5 Total 800-1800J | 1470nm, Pulse 7w 1s Ton, 5s Toff |
| 3、Tumour | 1470nm, Pulse 7w, 1s Ton, 5s Toff |
| 4、Tumour | 980m, Pulse 8w 1s Ton, 5s Toff |

|  |  |
| --- | --- |
| Dermatology | |
| 1、Cut | 980nm, CW 10W |
| 2、Coagulation | 980nm, CW 6w |
| 3、Moles Removal | 980nm, CW 8w |
| 4、Warts Removal | 980nm, CW 10w |
| 5、Nail Fungus | 980nm, Pulse, 6w, 1s Ton, 1s Toff |
| 6、Spider Vein 0.5mm | 980nm, Pulse, 8w, 1s Ton, 5s Toff |
| 7、Spider Vein 1mm | 980nm, Pulse, 8w, 1s Ton, 5s Toff |
| 8、Spider Vein 1.5mm | 980nm, Pulse, 8w, 1s Ton, 5s Toff |

|  |  |
| --- | --- |
| Liposuction | |
| 1、Liposuction Under Eye | 1470nm, 3w CW |
| 2、Liposuction Chin | 1470nm 6w CW |
| 3、Liposuction Arm | 1470nm 8w CW |
| 4、Liposuction Abdomen | 1470nm 12 CW |
| 5、Liposuction Buttock | 1470nm 10w CW |
| 6、Liposuction Thigh | 1470nm 8w CW |
| Gynecomastia | 1470nm 8w CW |

|  |  |
| --- | --- |
| 1、Therapy | |
| 2、Orthopedie Pain | 980nm 8w pulse, 50ms on, 50ms off |
| 3、Spine Pain | 980nm 8w pulse 50ms on, 50ms off |
| 4、Frozen Shoulder | 980nm 10w pulse 50ms on, 50ms off |
| 5、Plantar Facilities | 980nm 6w pulse 50ms on, 50ms off |
| 6、Pain Management | 980nm 6w pulse 50ms on, 50ms off |
| 7、Would Healing | 635nm CW |
| 8、Low Laser Therapy | 635nm CW |
| 9、Biostimulation | 635nm CW |